## Instructions
SUPPLEMENTARY MATERIAL 2. Uncertainty calculation for the mass
concentration of sirolimus in blood using the CRM ERMâ DA111a to study the bias
UNCERTAINTY ESTIMATION SPREADSHEETS
Version 1 (Raúl Rigo-Bonnin, (0034) 93 260 75 43, raulr@bellvitgehospital.cat)
IMPORTANT NOTE
These spreadsheets have been developed with all care and attention. It does not come with a guarantee for being error-free in the workings or in the interpretation of the uncertainty calculation.
Please bring any errors or questions to the attention of Raúl Rigo
Phone: (0034) 93 260 75 43, raulr@bellvitgehospital.cat
Uncertainty estimation
The aim is to facilitate clinical laboratories the estimation of measurement uncertainty. This Supplementary material contains five different spreadsheeds. The first three spreadsheets allow to estimate the main measurement uncertainty sources (ucal, uRw and ub). The fourth have been generated to calculate the combined standard uncertainty as well as the expanded uncertainty, and to know if the values obtained accomplish with the measurement uncertainty requirements previously selected by the laboratory. The fifth describe how to specify the measurand.
Calculation Process
IMPORTANT NOTE: The process is OPERATOR DEPENDENT. Skill and attention to detail is important for
producing valid results. If in doubt seek expert advice and/or another operator for validation.
I am happy to be consulted for advice.
Data Entry
ENTRIES INTO THE SPREADSHEETS MUST ONLY BE MADE IN BLUE BOXES
1. | In the spreadshet "ucal", add the biological quantity, measurement system and measurement units into their respective boxes.
2. | In the spreadsheet "ucal", add the assigned values of the end-user calibrator, the expanded uncertainties (in units) provided by the manufacturer, and the coverage factors used by the manufacturer into their respective boxes.
3. | In the spreadsheet "uRw", add the data obtained by your laboratory. That is, for each IQC level, add the IQC mean values, the standard deviations, and the number of IQC values into their respective boxes.
4. | In the spreadsheet "ub", add the data obtained by your laboratory after perform the bias study into their respective boxes. That is, the mean value, standard deviation and the number of CRM replicates. Also, add the CRM assigned value and its uncertainty (in units) facilitated by the CRM's manufacturer in their corresponding boxes.
5. | In the spreadsheet "COMBINED & EXPANDED UNCERTAINTY", add the maximum permissible percent relative root mean square of measurement error obtained from RiliBÄK (if it exists) or, failing that, add the maximum allowable coefficient of variation (in %) and the maximum allowable percent relative bias (in %) in their corresponding boxes.
NOTE: It is important that if the box corresponding to the RiliBAK requirements is not used, it must not contain any value (that is, it must remains empty).

## ucal
Biological quantity: | B--Sirolimus; subst.c.(ERM® DA111a; Acquity® UPLC®-TQD®)
Measurement system: | Acquity® UPLC®-TQD®
Measured values studied for the estimation of measurement uncertainty (IQC values): | 4.31 | 7.94 | 12.4
Measurement unit: | µg/L
Uncertainty related to the assigned value of the end-user calibrator
Measured value studied (x): | 4.31
Assigned value of the end-user calibrator (xcal): | 3.48
Expanded uncertainty (in units) provided by the manufacturer (Ucal): | 0.0984
Coverage factor used by the manufacturer (k): | 2
ucal = | 0.0492
Measured value studied (x): | 7.94
Assigned value of the end-user calibrator (xcal): | 6.43
Expanded uncertainty (in units) provided by the manufacturer (Ucal): | 0.1584
Coverage factor used by the manufacturer (k): | 2
ucal = | 0.0792
Measured value studied (x): | 12.4
Assigned value of the end-user calibrator (xcal): | 13.4
Expanded uncertainty (in units) provided by the manufacturer (Ucal): | 0.2834
Coverage factor used by the manufacturer (k): | 2
ucal = | 0.1417

## uRw
Biological quantity: | B--Sirolimus; subst.c.(ERM® DA111a; Acquity® UPLC®-TQD®)
Measurement system: | Acquity® UPLC®-TQD®
Measured values studied for the estimation of measurement uncertainty (IQC mean values): | 4.31 | 7.94 | 12.4
Measurement unit: | µg/L
Uncertainty related to the long-term intermediate imprecision
Level 1
Measured value studied (x): | 4.31
IQC mean value obtained for the measuring system (`x ): | 4.31
Standard deviation obtained for the measuring system (s): | 0.33
Number of IQC values processed for the measuring system (n): | 210
uRw = | 0.33
Level 2
Measured value studied (x): | 7.94
IQC mean value obtained for the measuring system (`x ): | 7.94
Standard deviation obtained for the measuring system (s): | 0.4838
Number of IQC values processed for the measuring system (n): | 210
uRw = | 0.4838
Level 3
Measured value studied (x): | 12.4
IQC mean value obtained for the measuring system (`x ): | 12.4
Standard deviation obtained for the measuring system (s): | 0.6282
Number of IQC values processed for the measuring system (n): | 38
uRw = | 0.6282

## ub
Biological quantity: | B--Sirolimus; subst.c.(ERM® DA111a; Acquity® UPLC®-TQD®)
Measurement system: | Acquity® UPLC®-TQD®
Measurement unit: | µg/L
Uncertainty related to the bias associated with the calibration process (using a CRM)
CRM mean value obtained after bias study (`x): | 11.27
Standard deviation obtained after bias study (s): | 0.77
Total number of CRM values processed (n): | 54
CRM assigned value (m): | 10.64
Uncertainty associated with the CRM assigned value (um): | 0.301
Correction factor to apply (cf): | 0.9441
Bias (b): | 0.63
um = ub = ucf = | 0.301
Is the bias significant ? (½b½> 2·ub): | YES

## COMBINED & EXPANDED UNCERTAINTY
Biological quantity: | B--Sirolimus; subst.c.(ERM® DA111a; Acquity® UPLC®-TQD®)
Measurement system: | Acquity® UPLC®-TQD®
Measured values studied for the estimation of measurement uncertainty (IQC values): | 4.31 | 7.94 | 12.4
Measurement unit: | µg/L
Maximum permissible percent relative root mean square of measurement error (%Drel(max)) obtained from RiliBÄK (if applicable):
Maximum allowable coefficient of variation (CVmax): | 15 | 15 | 15
Maximum allowable percent relative bias (%brel(max)): | 15 | 15 | 15
Maximum allowable percent relative expanded uncertainty (%Urel(max)): | 21.2132 | 21.2132 | 21.2132
Maximum allowable expanded uncertainty (Umax): | 0.91429 | 1.6843 | 2.6304
Combined standard uncertainty
Measured value studied (x): | 4.31 | 7.94 | 12.4
Uncertainty related to the assigned value of the end-user calibrator (ucal): | 0.0492 | 0.0792 | 0.1417
Uncertainty related to the long-term intermediate imprecision (uRw): | 0.33 | 0.4838 | 0.6282
Uncertainty associated with the bias (ub): | 0.301 | 0.301 | 0.301
uc = | 0.44936 | 0.57527 | 0.71086
%urel(c) = | 10.4259 | 7.2452 | 5.7327
Expanded uncertainty
Measured value studied (x): | 4.31 | 7.94 | 12.4
Combined standard uncertainty (uc): | 0.44936 | 0.57527 | 0.71086
Percent relative combined standard uncertainty (%rel(c)): | 10.4259 | 7.2452 | 5.7327
Coverage factor (k): | 2 | 2 | 2
Measured value studied (x) = | 4.31 | 7.94 | 12.4
U= | 0.89871 | 1.1505 | 1.4217
%Urel = | 20.8518 | 14.4904 | 11.4654
Is the expanded uncertainty obtained acceptable ? (U£ Umax): | YES | YES | YES
Is the percent relative expanded uncertainty obtained acceptable ? (%Urel £ %Urel(max)): | YES | YES | YES

## Specification of measurand
The measurand can be defined as:
The substance concentration (in nmol/L) of sirolimus in blood measured according to an in-house measurement procedure
using an Acquity® UPLC®-TQD® measuring system (Waters, Milford, MA, USA).
The results are traceable to assigned values of CRM ERMâ DA111a
Furthermore, the quantity can also be described using the IUPAC-IFCC nomenclature as:
B¾Sirolimus; subst.c.(ERMâ DA111a; Acquity® UPLC®-TQD®)