## Instructions
SUPPLEMENTARY MATERIAL 3. Uncertainty calculation for the mass
concentration of albumin in serum using IQC materials to study the bias
UNCERTAINTY ESTIMATION SPREADSHEETS
Version 1 (Raúl Rigo-Bonnin, (0034) 93 260 75 43, raulr@bellvitgehospital.cat)
IMPORTANT NOTE
These spreadsheets have been developed with all care and attention. It does not come with a guarantee for being error-free in the workings or in the interpretation of the uncertainty calculation.
Please bring any errors or questions to the attention of Raúl Rigo-Bonnin
Phone: (0034) 93 260 75 43, raulr@bellvitgehospital.cat
Uncertainty estimation
The aim is to facilitate clinical laboratories the estimation of measurement uncertainty. This Supplementary material contains five different spreadsheeds. The first three spreadsheets allow to estimate the main measurement uncertainty sources (ucal, uRw and ub). The fourth have been generated to calculate the combined standard uncertainty as well as the expanded uncertainty, and to know if the values obtained accomplish with the measurement uncertainty requirements previously selected by the laboratory. The fifth describe how to specify the measurand.
Calculation Process
IMPORTANT NOTE: The process is OPERATOR DEPENDENT. Skill and attention to detail is important for
producing valid results. If in doubt seek expert advice and/or another operator for validation.
I am happy to be consulted for advice.
Data Entry
ENTRIES INTO THE SPREADSHEETS MUST ONLY BE MADE IN BLUE BOXES
1. | In the spreadshet "ucal", add the biological quantity, measuring systems and measurement units into their respective boxes.
2. | In the spreadsheet "ucal", add the assigned values of the end-user calibrator, the expanded uncertainties (in units) provided by the manufacturer, and the coverage factors used by the manufacturer into their respective boxes.
3. | In the spreadsheet "uRw", add the data obtained by your laboratory. That is, for every measuring system used and for each IQC level, add the IQC mean values, the standard deviations, and the number of IQC values into their respective boxes.
4. | In the spreadsheet "ub", for every measurement system used and for each IQC level, add the robust standard deviation, the number and the mean of laboratories participating in an inter-laboratory quality control scheme in their corresponding boxes.
5. | In the spreadsheet "COMBINED & EXPANDED UNCERTAINTY", add the maximum permissible percent relative root mean square of measurement error obtained from RiliBÄK (if it exists) or, failing that, add the maximum allowable coefficient of variation (in %) and the maximum allowable percent relative bias (in %) in their corresponding boxes.
NOTE: It is important that if the box corresponding to the RiliBAK requirements is not used, it must not contain any value (that is, it must remains empty).

## ucal
Biological quantity: | Srm--Albumin; mass c.(CRM 470; Cobas 8000 or Cobas 6000)
Measuring system 1: | Cobas 8000-1
Measuring system 2: | Cobas 8000-2
Measuring system 3: | Cobas 8000-3
Measuring system 4: | Cobas 6000
Measured values studied (IQC mean values): | 27.1165 | 47.6421
Measurement unit: | g/L
Uncertainty related to the assigned value of the end-user calibration material
Measured value studied (x): | 27.1165
Assigned value of the end-user calibrator (xcal): | 35.2
Expanded uncertainty (in units) provided by the manufacturer (Ucal): | 0.298
Coverage factor used by the manufacturer (k): | 2
ucal = | 0.149
Measured value studied (x): | 47.6421
Assigned value of the end-user calibrator (xcal): | 35.2
Expanded uncertainty (in units) provided by the manufacturer (Ucal): | 0.298
Coverage factor used by the manufacturer (k): | 2
ucal = | 0.149

## uRw
Biological quantity: | Srm--Albumin; mass c.(CRM 470; Cobas 8000 or Cobas 6000)
Measuring system 1: | Cobas 8000-1
Measuring system 2: | Cobas 8000-2
Measuring system 3: | Cobas 8000-3
Measuring system 4: | Cobas 6000
Measured values studied for the estimation of measurement uncertainty (IQC mean values): | 27.1165 | 47.6421
Measurement unit: | g/L
Uncertainty related to the long-term intermediate imprecision
Level 1
Measured value studied (x): | 27.1165
Measuring system (i) | Cobas 8000-1 | Cobas 8000-2 | Cobas 8000-3 | Cobas 6000
IQC mean value obtained for measuring system i (`xi ): | 27.57 | 27.62 | 26.42 | 26.91
Standard deviation obtained for measuring system i (si): | 0.62 | 0.56 | 0.52 | 0.54
Number of IQC values processed for measuring system i (ni): | 186 | 131 | 171 | 151
uRw = | 0.75372
Level 2
Measured value studied (x): | 47.6421
Measuring system (i) | Cobas 8000-1 | Cobas 8000-2 | Cobas 8000-3 | Cobas 6000
IQC mean value obtained for measuring system i (`xi ): | 48.57 | 48.77 | 46.76 | 46.54
Standard deviation obtained for measuring system i (si): | 0.67 | 0.52 | 0.53 | 0.73
Number of IQC values processedfor measuring system i (ni): | 175 | 124 | 164 | 143
uRw = | 1.1801

## ub
Biological quantity: | Srm--Albumin; mass c.(CRM 470; Cobas 8000 or Cobas 6000)
Measuring system 1: | Cobas 8000-1
Measuring system 2: | Cobas 8000-2
Measuring system 3: | Cobas 8000-3
Measuring system 4: | Cobas 6000
Measured values studied for the estimation of measurement uncertainty (IQC mean values): | 27.1165 | 47.6421
Measurement unit: | g/L
Uncertainty related to the bias using IQC materials
Level 1
Measured value studied (x): | 27.1165
Measuring system (k) | Cobas 8000-1 | Cobas 8000-2 | Cobas 8000-3 | Cobas 6000
IQC mean value obtained for the IQC material level 1 on measuring system k (`x1,k): | 27.57 | 27.62 | 26.42 | 26.91
Standard deviation obtained for the IQC material level 1 on measuring system k (s1,k): | 0.62 | 0.56 | 0.52 | 0.54
Number of replicate measurements using the IQC material level 1 on measuring system k (n1,k): | 186 | 131 | 171 | 151
Robust peer-group standard deviation for the IQC material level 1 on measuring system k (sLabs 1,k): | 0.943 | 0.94 | 0.94 | 1.01
Number of peer-group laboratories participating in the IQC material level 1 on measuring system k (q1,k): | 168 | 168 | 168 | 113
Mean of arithmetic means of peer-group laboratories participating in an internal-external program for the IQC material level 1 on measuring system k (m1,k ): | 26.32 | 26.32 | 26.32 | 25.2
Bias using the IQC material level 1 on measuring system k (b1,k ): | 1.25 | 1.3 | 0.1 | 1.71
Uncertainty associated with the mean of peer-group laboratories participating in an internal-external program for the IQC material level 1 on measuring system k (um ,1,k ): | 0.090943 | 0.090653 | 0.090653 | 0.11877
Level 2
Measured value studied (x): | 47.6421
Measuring system (k) | Cobas 8000-1 | Cobas 8000-2 | Cobas 8000-3 | Cobas 6000
IQC mean value obtained for the IQC material level 1 on measuring system k (`x1,k): | 48.57 | 48.77 | 46.76 | 46.54
Standard deviation obtained for the IQC material level 1 on measuring system k (s1,k): | 0.67 | 0.52 | 0.53 | 0.73
Number of replicate measurements using the IQC material level 1 on measuring system k (n1,k): | 175 | 124 | 164 | 143
Robust peer-group standard deviation for the IQC material level 1 on measuring system k (sLabs 1,k): | 1.16 | 1.2 | 1.09 | 1.19
Number of peer-group laboratories participating in the IQC material level 1 on measuring system k (q1,k): | 145 | 146 | 143 | 125
Mean of arithmetic means of peer-group laboratories participating in an internal-external program for the IQC material level 1 on measuring system k (m1,k ): | 46.24 | 46.24 | 46.24 | 45.92
Bias using the IQC material level 1 on measuring system k (b1,k ): | 2.33 | 2.53 | 0.52 | 0.62
Uncertainty associated with the mean of peer-group laboratories participating in an internal-external program for the IQC material level 1 on measuring system k (um ,1,k ): | 0.12042 | 0.12414 | 0.11394 | 0.13305
`bw | 1.2639
u`b | 0.81397
Is the bias significant ? (½`bw ½> 2 · u`b): | NO

## COMBINED & EXPANDED UNCERTAINTY
Biological quantity: | Srm--Albumin; mass c.(CRM 470; Cobas 8000 or Cobas 6000)
Measuring system 1: | Cobas 8000-1
Measuring system 2: | Cobas 8000-2
Measuring system 3: | Cobas 8000-3
Measuring system 4: | Cobas 6000
Measured values studied for the estimation of measurement uncertainty (IQC values): | 27.1165 | 47.6421
Measurement unit: | g/L
Maximum permissible percent relative root mean square of measurement error (%Drel(max)) obtained from RiliBÄK (if applicable): | 12.5 | 12.5
Maximum allowable coefficient of variation (CVmax):
Maximum allowable percent relative bias (%brel(max)):
Maximum allowable percent relative expanded uncertainty (%Urel(max)): | 12.5 | 12.5
Maximum allowable expanded uncertainty (Umax): | 3.3896 | 5.9553
Combined standard uncertainty
Measured value studied (x): | 27.1165 | 47.6421
Uncertainty related to the assigned value of the end-user calibrator (ucal): | 0.149 | 0.149
Uncertainty related to the long-term intermediate imprecision (uRw): | 0.75372 | 1.1801
Bias (b): | 0 | 0
uc = | 0.76831 | 1.1895
%urel(c) = | 2.8334 | 2.4967
Expanded uncertainty
Measured value studied (x): | 27.1165 | 47.6421
Combined standard uncertainty (uc): | 0.76831 | 1.1895
Percent relative combined standard uncertainty (%rel(c)): | 2.8334 | 2.4967
Coverage factor (k): | 2 | 2
Measured value studied (x) = | 27.1165 | 47.6421
U= | 1.5366 | 2.3789
%Urel = | 5.6667 | 4.9934
Is the expanded uncertainty obtained acceptable ? (U£ Umax): | YES | YES
Is the percent relative expanded uncertainty obtained acceptable ? (%Urel £ %Urel(max)): | YES | YES

## Specification of Measurand
The measurand can be defined as:
The mass concentration (in g/L) of albumin in serum measured using Cobas® 8000 and Cobas® 6000
measuring systems (Roche Diagnostics, Risch-Rotkreuz, Switzerland).
The results are traceable to the certified reference material CRM 470 prepared from the international conventional
standard USNRP 12-0575C
Furthermore, the quantity can also be described using the IUPAC-IFCC nomenclature as:
Srm¾Albumin; mass c.(CRM 470; Cobas 8000 or Cobas 6000)